- POLICY

  - General instructions for laboratory personnel during a scheduled or unscheduled downtime to allow for continued delivery of patient results.

  - The following procedure will be followed in the event of downtime involving:

    - LIS failure

      - IHIS failure (IHIS and/or LIS DI middleware interfaces)

      - WAM failure (Hematology LIS middleware interface)

    - Instrument failure

    - Power failure

    - Water failure

    - Bio-Rad Unity Software failure

    - Any other testing delays

- PURPOSE OF DOCUMENT

  - To ensure reporting of patient results occurs in a prompt and efficient manner during partial or complete system downtime until recovery of the system.

- SCOPE OF DOCUMENT

  - This document applies to all personnel at the Morehouse Tower Laboratory.

- RESPONSIBILITY

  - The lead/manager is responsible for maintaining the document and ensuring biennial review. The laboratory medical director is responsible for establishing and approving all changes before activating document, and reviewing at least biennial.

- PROCESS

  - Refer to Admin-148 Downtime Procedure

    - Locate downtime binder with downtime labels and a limited amount of downtime requisitions.  (Requisitions can be purchased through RR Donnelley by the patient care unit/department.)

  - Document downtime on MMP-F45 Instrument Downtime Form

  - Notification of Testing Delays: The following process is used for any testing delays, including those due to unavailable reagents, extended instrument / equipment down times, or significant quality failures (e.g. proficiency testing, quality controls, calibrations).

    - Testing personnel: notify Lead Technologist immediately when problems are identified which will delay testing beyond published turn-around time. In the absence of the Lead Technologist, notify the Manager or Administrative Director.

    - Lead Technologist or designee will notify manager and division director(s) of delays immediately. They will also the Nurse Managers for Chemo and other clinics serviced by the Martha Morehouse Laboratory.

    - Lead Technolgist, Manager, and Division Director:

      - Determine corrective action to be taken.

      - Hold specimens until testing is available again, OR

      - Send specimens to other OSU laboratory.

      - Determine whether medical staff (internal and external) should be notified.

  - In the event of LIS failure, all stat reports will be generated from the instruments, properly labeled and delivered or faxed to the appropriate nursing stations in the Morehouse Tower.

    - Report IHIS failure to the HelpDesk at 293-3861

      - Forward Help Desk incident ticket to LIS email group [pathologyLIS@osumc.edu].

      - If only live (production) IHIS is down, the Downtime (shadow) version of IHIS may be available. This is read only but allows searching and reviewing patient information similar to the live environment.

5.4.1.b..1. IHIS Downtime can be accessed by clicking the Start Menu > All Programs > IHIS

Downtime (folder) > IHIS Downtime. Log in with your IHIS user name and password.

      - Complete IHIS failure:

5.4.1.c..1. At 30 minutes, fax and/or tube all stat results.

5.4.1.c..2. At 45 minutes, notify the laboratory manager, administrative director, and

division(s) directors.

      - IHIS DI Middleware (interface) downtime:

5.4.1.d..1.  At 30 minutes, begin manually entering results into IHIS.

5.4.1.d..2.  At 45 minutes, notify the laboratory manager, administrative director, and division(s)

directors.

5.4.1.d.3. Note: Specimens can still be received in IHIS but analyzers will not recognize barcode

information. All orders require manually programming on the analyzer using the master           patient index (MPI). The MPI is a unique barcode number. Program the analyzer using two patient identifiers (patient legal name and MRN), the MPI, and each test requested.

    - Report WAM downtime to CCL Hematology at 293-3444.

      - Refer to 5 L Heme-33 Hematology Downtime Policy.

      - At 30 minutes, begin manually entering results into IHIS.

      - At 45 minutes, notify the laboratory manager, administrative director, and division(s) directors.

    - Set analyzer(s) to print all patient results.

a.	Hematology – XN2000

5.4.3.a.1. Log-in as admin

5.4.3.a.2. Menu>IPU setting>Auto Output

5.4.3.a.3. Under “Auto Output Destination and Output Conditions”

( Check the small box to the left of GP

5.4.3.a.3.1.	Note: When downtime resolves, uncheck small box to the left of GP. (The small box to the left of HC should be the only box checked.)

5.4.3.a.4. Select Apply

5.4.3.a.5. Select OK

b.  Chemistry - DXC700AU

5.4.3.b.1. Must be in standby.

5.4.3.b.2. From Menu, select Config., select Format, select List Format

5.4.3.b.3. Select Edit (F1)

5.4.3.b.4. Select Realtime List (F5)

5.4.3.b.5. Under Patient select “Patient Report”, select OK

5.4.3.b.5.1.	Note: When downtime resolves, choose “none” under Patient.

5.4.3.b.6. Select Save (F1)

- Coagulation - STA Compact MAX

5.4.3.c.1. Select SYSTEM>Select Online printing (verify check mark remains by Online printing)

5.4.3.c.1.1: Ensure printer USB cord is plugged into the Compact Max.

5.4.3.c.1.2: When downtime resolves, uncheck Online Printing.

- Urinalysis - Clinitek Status +

5.4.3.d.1. Home Screen>Instrument Setup>Instrument Settings> Printer Settings

5.4.3.d.2. Scroll through screens to select Enabled>Next>Internal Printer>Done

5.4.3.d.2.1. 	Note: When downtime resolves, choose “External Printer”.

    - During the downtime refer to Admin-148 Downtime Procedure for receiving, labeling, linking and test resulting based on the scenario of the downtime.

    - Call critical results to the appropriate floor and write all necessary critical call information on the printout to document the call.  The callback information must be entered into the LIS when the system is functional.

    - Fax reports to the appropriate clinic or floor as they are generated. Verify appropriate fax numbers each downtime.

  - In the event of instrument failure or water system failure, which has been determined to exceed one hour, all stat specimens will be delivered, by courier, to another OSU laboratory where they will be resulted in the LIS.  The Lead, Manager or Director will determine which laboratory is most appropriate.

  - In the event of Bio-Rad QC Software failure, manually compare results to the current mean and SD backup spreadsheet at:  L:\Shared\Pathology\MMMP\ QC Current Means Unity to compare QC results with current Unity Real Time Means/SD. Verify QC results fall within the acceptable 2SD/3SD range prior to performing patient testing.

  - In the event of power failure that would render any instrumentation non-functional, a cease of operations takes place and all specimens and reagents (if necessary) will be transported to another OSU laboratory.  The reagents will be stored according to manufacturer guidelines on an individual basis.

    - NOTE: Analyzers, refrigerators and freezers are on back-up power.

    - In the event of power failure, where lab is still operational due to backup power, a centrifuge may need moved to another counter to be plugged in to back up power.

    - A total, planned power outage of back up power (red outlets) of any length, may require removal of on-board refrigerated reagents from the chemistry instruments.  Reagents should be capped and reagents wheels removed from the analyzer and placed into the refrigerators to maintain acceptable stability.

    - Recovering from a total, lengthy power outage (including back-up power)

Note: Not necessary for monthly generator testing, however, if there are issues after generator testing, these things can be checked.)

      - Verify Medica water systems are on and running.  If not, turn on and/or press Process button.

      - Turn on all instruments.

        - XN2000

          - Acknowledge any errors on the RU-20 if applicable.

          - Note: RU-20 may take longer to recovery from power failure.

        - DXC700AU

          - If there is a power failure, the analyzer immediately stops operation. Power to the incubator and reagent refrigerator is also turned off.

          - Press the RESET button (white) to turn on the main power, and then wait 5 seconds.

          - Press the ON button (green). The lamp turns on and the software loads. The system displays a dialog to confirm retrieving the database.

          - The system is in WARM UP mode for 1.5 hours. After the required 20-minute lamp warm up time, wait until the temperature of the cuvette wheel is 37 °C, and then select MAINT. > Analyzer Maintenance > Maintenance. Select Standby [F4] to return to STANDBY mode.

          - Note: If the power failure occurred during patient testing you must perform a W1. Sample can remain in the sample probe, and reagents can remain in the cuvettes. Perform a W1 to clean the sample probe and cuvettes after you restart the system.

          - Verify the Home Analyzer Status screen is all green (acceptable limits)

          - Load reagent wheels if applicable.

          - Run QC to confirm the reagent integrity before resuming analysis.

        - STA Compact MAX

          - Launch the program software.

      - Turn on all computers.

        - Unity Real Time computer

          - Allow the Unity program to auto-launch completely after turning on, do not attempt to open Unity Connect first, then stop and start the data manager.

      - Restart the heat block by pressing the START button.

      - Log into Isensix to review humidity and temperature sensors.  Document all events.

- RELATED DOCUMENTS

  - Refer to QPulse System or Document Detail Report for related Laboratory Policies, Procedures, and Master Forms